## Page1
Certificate of Calibration
Arbin Instruments Calibration Laboratory
Certificate Number: 215183-X2-11132020
| Model Number | PCA425554 | Customer
| Serial Number | 215183-X2 | TM Technology JSC No 6/21, Lane 63/33, Le Duc Tho Street My Dinh 2 Ward Nam Tu Liem District, Hanoi Vietnam
| Manufacturer | Arbin Instruments
| Description | Voltage Range 5V, Current Range 1A/0.1A
| Certificate Issue Date | November 14th, 2020 | Location of Calibration
| Date of Calibration | November 13th, 2020 | Arbin Instruments Calibration Laboratory 762 Peach Creek Cut Off Road College Station, TX 77845
| Calibration Interval | 12 MONTHS
| Calibration Due Date | November 13th, 2021
| Date Type | As Found/ As Left/ Found-Left
| Temperature/RH | 23°C ± 1°C / 50% ± 20%
| Procedure | Arbin Calibration Manual
This calibration is traceable to the International System of Units (SI) through the National Institute of Standards and Technology (NIST) and/or recognized national or international standards laboratories. This certificate applies only to the item identified and shall not be reproduced without prior written approval of Arbin Instruments Laboratory. Certificates of Calibration without signature are invalid. This laboratory is accredited in accordance with the recognized International Standard ISO/IEC 17025. This accreditation demonstrates technical competence for a defined scope and the operation of a laboratory quality management system (refer to joint ISOILAC-IAD Communiqué dated April 2017) in electrical calibration. The auto calibration process requires measuring the output voltage and current 5 times, then obtains the upper and lower limit based on the readings from the auto calibration product named ACB. After that, the multimeters measure the output values and compare them with the ACB measurements to obtain accuracy. In order to pass the calibration, the accuracy of measurements should be within ±40 ppm. For example, if the current output value of the current comparator measured by multimeters is 1A, then the allowed error range is ±1A*2*40ppm/1000000=80µA. It means for a 1A current in the 1A current range, the accepted range is 0.99992A to 1.00008A. If the average value of current measurement from the ACB is within this range, then the ACB product will pass the calibration, otherwise the ACB has failed calibration. Currently the measurement of uncertainty is focused on accuracy, but precision is not taken into account in the process of evaluating pass and fail criteria.
Arbin Instruments
762 Peach Creek Cut Off Road
College Station, TX 77845
Phone (979) 690-2751
Fax (979) 690-2761 |  |  |  | Page1
www.arbin.com

## Page2
Certificate of Calibration
Arbin Instruments Calibration Laboratory
Certificate Number: 215183-X2-11132020
Accreditation #103832
The Data Type in this certificate must be interpreted as: As Found Calibration data collected before the unit is adjusted and / or repaired. As Left Calibration data collected after the unit has been adjusted and / or repaired. Found-Left Calibration data collected without any adjustment and / or repaired performed. This Certificate may contain data that is not covered by the Scope of Accreditation. Unaccredited material, where applicable is indicated by an asterisk (*), or confined to clearly marked sections. Functional (Pass/Fail) testes are not accredited.
Allen Livingston, Calibration Engineer | Mu Li, QC Approval
Arbin Instruments
762 Peach Creek Cut Off Road
College Station, TX 77845
Phone (979) 690-2751
Fax (979) 690-2761 |  | Page2
www.arbin.com

## Page3
Certificate of Calibration
Arbin Instruments Calibration Laboratory
Certificate Number: 215183-X2-11132020
Standards Used
Model Description | Equipment ID/SN | Cal Date | Cal Due Date | Certification No.
8508A Fluke Digital Multimeter 8½ Digit | 282268384 | 2020-10-19 | 2022-10-19 | EVL582676
8508A Fluke Digital Multimeter 8½ Digit | 297968830 | 2020-10-21 | 2022-10-21 | EVL583021
9331/1 M.I. 1 Ohm Standard Air Resistor | 1102801 | 2019-03-23 | 2021-03-23 | C1180310
6011D/300 M.I. DC Current Comparator Range Extender | 1102693 | 2019-03-25 | 2021-03-25 | C1180311
Arbin Instruments
762 Peach Creek Cut Off Road
College Station, TX 77845
Phone (979) 690-2751
Fax (979) 690-2761 |  |  |  |  | Page3
www.arbin.com

## Result Page
Certificate of Calibration
Arbin Instruments Calibration Laboratory
Certificate Number: 215183-X2-11132020
Calibration Result
Parameter | Range | Nominal | Lower Limit | As Found | As Left | Upper Limit | Result | Uncertainty
Tested | [6 empty] | Pass/Fail
Voltage | 5V | 0.12346 | 0.12346 | 0.12346 | 0.12346 | 0.12346 | Passed | 5nV
|  | 0.12346 | 0.12346 | 0.12346 | 0.12346 | 0.12346 | Passed | 5nV
Arbin Instruments
762 Peach Creek Cut Off Road
College Station, TX 77845
Phone (979) 690-2751
Fax (979) 690-2761 | [8 empty] | Page4
www.arbin.com